## Slide 1: TESA UPC / UPD Gauge block comparator
PMI department
February 2013

## Slide 2: Summary
Gauge Blocks Introduction
TESA UPC Gauge Block Comparator
        - For comparative measurements
TESA UPD Gauge Block Comparator
- For comparative and direct measurements
TESA UPT Temperature Device
TESA UP The Software Programme
Marketing Delivery Programme / Maintenance / Investments
TESA UPC / UPD Configuration
TESA UPC / UPC System Components

## Slide 3
Welcome to the world of gauge block calibration
TESA Technology offers two instruments for two different ways of calibration
Gauge Blocks - Introduction
I was the first
Made by Brown & Sharpe for Mr. C.E. Johansson

## Slide 4
Gauge blocks:
The most significant material measures in dimensional metrology
Gauge Blocks - Introduction
Their periodic calibration ensures that they remain permanently reliable

## Slide 5
as per ISO 3650:1998
Ln = nominal length
te = limit deviation at any point proceeding from the nominal length
lc = central length
fu & fo = variation v
tv = tolerance of the variation v
Gauge Blocks - Introduction

## Slide 6
Gauge Blocks - Introduction
Tracing gauge blocks (material measures) to the unit of length (meter)

## Slide 7
Acquisition of the gauge length with use of two probes.
Configuration layout
TESA UPC / UPD - Configuration
Measuring support
Measuring table
Upper probe
Lower probe
Electornique module
Pneumatic probe retraction (only for UPC)
Reference gauge block
Calibrated gauge block

## Slide 8
TESA UPC / UPD - Configuration
The economic and metrologic advantages
of a dual probes measuring system:
The two probes in «sum measurement»
mode ensure a correct measurement.
No form and position errors influences.
Especially for small thickness gauge blocks.
No surface apposition makes the measuring cycles faster.
Templates systems measuring lc point on the reference block and
lc plus 4 other points on each corner of the block being measured

## Slide 9
Measuring gauge block of same nominal length.
Application ranges from 0,5 up to 100 mm (up to 500 mm in specialty.)
TESA high-precision inductive probes.
Template system for gauge block positionning.
Computer-aided data processing software available
Ultra-precise temperature device can be integrated.
Transfer on-line all length (and optionally the temperature).
Performs calibrations by meeting ISO standards as well as EA guidelines.
Greater accuracy execution available with a calibration certificate.
For Comparative Measurement
TESA UPC – Gauge Block Comparator

## Slide 10
TESA UPC – Gauge Block Comparator
TESA UPC execution ref. 05930014

## Slide 11
Errors of measurement
Repeatability limit
Uncertainty of measurement
• Execution for greater accuracy
• Standard execution
*10,015 µm
*2U = ± (0,05 + 0,5 x L) µm
*10,025 µm
( L in m )
*2U = ± (0,10 + 1,0 x L) µm
TESA UPC – Gauge Block Comparator
( L in m )
*1 Without influence of the temperature
*2 Condition involves the use of reference standards whose uncertainty of measurement is:
U ≤ ± 0,015 µm for the comparator calibration U ≤ ± (0,02 + 0,2 x L) µm for the gauge block calibration

## Slide 12
TESA UPD – Gauge Block Comparator
For Direct and Comparative Measurement with a 25mm span
The flexible concept that provides distinctive metrological features with substantial savings at the end of it all

## Slide 13
TESA UPD – Gauge Block Comparator
For Direct and Comparative Measurement with a 25mm span
Measuring gauge block of same or different nominal length.
Application ranges from 0,5 up to 100 mm (up to 500 mm in speciality.)
TESA & HEIDENHAIN high-precision incrementale probes.
Template systems for gauge block positionning.
Computer-aided data software available.
Ultra-precise temperature device can be integrated.
Transfer on-line all lengths (and optionally the temperature).
Performs calibrations by meeting ISO standards as well as EA guidelines.

## Slide 14
TESA UPD – Gauge Block Comparator
TESA UPC execution ref. S59300102

## Slide 15
TESA UPD – Gauge Block Comparator
The advantages that make it leader and unrivaled:
The Time:
The Costs:
Number of probing points is thus reduced drastically.
Measured value immediately displayed after the probe insert has come into contact.
Number of calibrated gauge blocks increased with a single probed reference gauge
Allows to calibrate over 90% of a 122-pieces gauge block set with the same reference standard.
Consequently, it allows to reduce the reference set compared to the UPC.
Reduce significantly the cost of purchase or maintenance on the references.
No need to buy specific reference for an unusual nominal length (e.g. 17.6mm)

## Slide 16
Comparative measurement
Comparison of gauge blocks when same nominal length
Direct measurement
Comparison with a variation in nominal length of 25mm
TESA UPD – Gauge Block Comparator
for smaller systematic errors

## Slide 17
Upper HEIDENHAIN high-precision axial probes
TESA UPD – Gauge Block Comparator
11111100110
- 25mm measuring span
- Opto-electronic measuring system
- ZERODUR glass ceramic scale with incremental divission
- ABBE’s principle fully respected
- Accuracy ≤ 5 nm
- Motorised measuring bolt activation
- 0.000005 mm (5 nm) numerical interval

## Slide 18
TESA UPD – Gauge Block Comparator
Errors of measurement
Repeatability limit
Uncertainty of measurement
*10,015 µm
*2U = ± (0,05 + 0,5 x L) µm
*1 Without influence of the temperature
*2 Condition involves the use of reference standards whose uncertainty of measurement is:
U ≤ ± 0,015 µm for the comparator calibration U ≤ ± (0,02 + 0,2 x L) µm for the gauge block calibration
( L in m )

## Slide 19
… also used on measuring equipments such as horizontal measuring benches or height gauges.
TESA UPT – Temperature Device
For TESA gauge block comparators UPC and UPD
4 temperature sensors PT100 (4-wires type).
2 of them with clamps for the gauge blocks
Numerical interval: 0.001°C
Calibrated for the measuring range from 19°C up to 24°C
Uncertainty of measurement: U = +/- 0.03°C
Delivered with SCS calibraiton certificate

## Slide 20
TESA UP – Software Programme
For TESA gauge block comparators UPC and UPD
On-line processing of length and temperature values.
Automatic execution of all relevant corrections.
Large possibilities for a self-creation of certificates and calibration modes.
Result outputs according to ISO 3650:1998, to 6 other standards and to a personnalized one.
Available in 11 languages.
Compatible with Win 7 (32 bits) and lower.
Compatible for comparators from competitors.
Delivered with protective dongle (hard-key) in USB

## Slide 21
Dual Template: a patented pending system for TESA UPD and optionnaly for UPC
Main target: to protect the expensive reference gauge blocks.
TESA UPC / UPD – System Components
How: by reducing the travel of the reference gauge blocks of about 70%.
This contribute to lower the risks of damaging and wearing the the measuring faces.
This double protection blocks leads to significant cost savings by reducing needs for:
Recalibration
Restoring measuring faces, replacing worn or damaged referenced gauges
Life extension of the full reference sets.
Longer life of the pins of the table of the UPC or UPD.

## Slide 22
TESA UPC – System Components

## Slide 23
TESA UPD – System Components

## Slide 24
TESA UPC / UPD – Delivery Programmes
TESA UPC: 8 Ordering variations
TESA UPD: 3 Ordering variations

## Slide 25
TESA UPC / UPD – Delivery Programmes
Options (depending the delivery programme) and Accessories:

## Slide 26: – TESA offer the widest range of measuring equipments for gauge block calibration Worldwide. - TESA UPD Gauge Block Comparator is the only one of its class that has no competitor.
Precision – our identity
Thanks !
You have more questions. A team of Enginneers is at your dispositon at:
tesa.service@hexagonmetorlogy.com => for technical requests / questions
tesa.specialites@hexagonmetorlogy.com => for non standard products quotation
TESA UPC / UPD - Conclusion